InfuLife  |  Staff Competency Assessment

Elastomeric Infusion Pump — Instructions for Use  |  Reference: PMP-1100 Rev E & PMP-1103 Rev A  |  First Medical Source (FMS), LLC

| Staff Name: ___________________________ | Date: ___________________________ | Unit/Department: ___________________________ | Evaluator: ___________________________ |
| --- | --- | --- | --- |

Instructions: 12 questions drawn from InfuLife IFU (PMP-1100 Rev D). Select the single best answer (A–D) and write it on the Answer line. Passing score: 80% (10/12). Scores below 80% require remediation and retest. Answer Key is on the final page — for evaluator use only.

Assessment Questions

| Q1 Flow Sensitivity Your patient's InfuLife pump has the flow restrictor positioned 10 inches higher than body level. According to the IFU, what effect does raising the flow restrictor 10 inches above the reference (body) height have on the infusion flow rate? A. The flow rate decreases by approximately 6% B. The flow rate increases by approximately 6% C. There is no clinically significant change in flow rate D. The flow rate increases by approximately 9% Answer: ________ |  | Q2 Temperature The flow restrictor is exposed on top of the patient's blanket at room temperature (23°C/73°F) rather than against the patient's skin (31°C/88°F). Compared to optimal placement against the skin, what happens to delivery time when the restrictor is at room temperature? A. Delivery time increases by approximately 9% B. Delivery time decreases by approximately 9% (drug delivered faster than intended) C. Delivery time increases by approximately 10% D. There is no significant impact — only fill volume affects delivery time Answer: ________ |
| --- | --- | --- |
| Q3 Temperature A pharmacy-prepared InfuLife has been refrigerated overnight. The RN plans to begin infusion immediately after removal. What does the IFU require before starting an infusion on a device stored in the refrigerator? A. Allow it to warm to room temperature before use B. Begin immediately — cold storage slows the rate to a safer level C. Prime the tubing with warm saline first, then connect D. Refrigerated storage is not permitted; the device must be discarded Answer: ________ |  | Q4 Fill Calculation Physician orders a 270 mL InfuLife at 5 mL/h for a 48-hour infusion. What is the correct fill volume to prepare? A. 240 mL (infusion time × flow rate only) B. 236 mL (from table PMP-4320) C. 250 mL (infusion time × flow rate + 10 mL residual volume) D. 270 mL (use nominal fill volume regardless) Answer: ________ |
| Q5 Flow Sensitivity An InfuLife is prepared with 5% dextrose (D5W) instead of the usual 0.9% sodium chloride. How does using 5% dextrose affect delivery time compared to normal saline? A. Delivery time is approximately 10% shorter B. Delivery time is approximately 10% longer C. Delivery time is approximately 6% longer D. There is no difference — viscosity effects apply only to temperature Answer: ________ |  | Q6 Contraindications Which of the following is a listed contraindication for InfuLife use? A. Subcutaneous delivery of local anesthetics B. Epidural infusion with a specifically designed epidural catheter C. Intra-articular infusion of local anesthetics D. Postoperative pain management at surgical wound sites Answer: ________ |

Assessment Questions (continued)

| Q7 Safety / Alarms A patient asks if the InfuLife will sound an alarm if the infusion stops or is blocked. What is the correct response? A. Yes — a soft audible alarm activates if flow is interrupted for more than 15 minutes B. No — there is no alarm or alert when flow interruption occurs C. An indicator light changes color when flow is interrupted D. The volume scale on the outer cover will flash to indicate a blockage Answer: ________ |  | Q8 Filling Procedure What must be done with the clamp before dispensing drug solution into the reservoir during filling? A. Open the clamp to allow air to escape as fluid fills the reservoir B. Close the clamp during filling to prevent premature flow C. The clamp position does not matter during the filling step D. Clamp should be half-open to maintain back-pressure Answer: ________ |
| --- | --- | --- |
| Q9 Flow Sensitivity The nominal flow rates listed for InfuLife are based on which reference conditions? A. 5% dextrose at 37°C (body temperature), at 5 inches of head height B. 0.9% sodium chloride at 32°C, at 0 inches of head height per ISO 28620:2020 C. Sterile water at 23°C (room temperature), at 0 inches of head height D. 0.9% sodium chloride at 37°C, at 6 inches of head height Answer: ________ |  | Q10 Patient Population Why is InfuLife contraindicated for neonates, infants, and children weighing less than 20 kg? A. The flow rate is too high for small body mass and cannot be adjusted B. The device is too large for pediatric IV access sites C. Potential exposure to ethylene oxide (EO) residuals D. The elastomeric bladder generates excessive pressure for small patients Answer: ________ |
| Q11 Flow Sensitivity An RN fills a 270 mL InfuLife with only 240 mL — 30 mL below nominal. What effect does underfilling the reservoir below nominal volume generally have on flow rate? A. Lower fill volume generally results in a lower flow rate B. Lower fill volume generally results in a higher flow rate due to reduced back-pressure C. Lower fill volume has no effect — the elastomer compensates D. Lower fill volume increases delivery time by exactly 10% Answer: ________ |  | Q12 Filling Procedure After priming the tubing and before connecting to the patient access site, what must the RN verify? A. That the filter is wet and no bubbles are visible in the filter housing B. That the clamp on the tubing is closed before attaching to the injection site C. That the pump has been at room temperature for at least 2 hours D. That the outer cover scale reads the correct fill volume Answer: ________ |

| Total Questions: 12 | Number Correct: _______ / 12 | Score: _______ % |
| --- | --- | --- |
| Pass (≥ 80%) / Fail (< 80%) | Evaluator Signature: _______________________ |  |

Answer Key & Rationale

For evaluator use only. Do not distribute to candidates prior to assessment.

| Q1 | Ans: B | Per the IFU: an increase/decrease of 10 inches in head height will increase/decrease flow rate by 6%. Raising the restrictor increases hydrostatic pressure, speeding delivery. Height management is critical for accurate dosing. |
| --- | --- | --- |
| Q2 | Ans: B | The IFU states: when the device is used with restrictor at room temperature (23°C/73°F), delivery time will decrease approximately 9%. The drug is infused faster than intended. Always position the restrictor against or close to the patient's skin. |
| Q3 | Ans: A | The IFU explicitly states: if the device is stored in refrigerator, allow it to warm to room temperature before use. Storage of a filled device beyond 8 hours before starting may also result in a 10% longer delivery time. Freezer storage is prohibited. |
| Q4 | Ans: B | 48h × 5 mL/h + residual = 245 mL, which exceeds the ±10 mL threshold. Per IFU, refer to overfill/underfill chart PMP-4320. Column C270R0050, row 48 hrs = 236 mL fill volume. |
| Q5 | Ans: B | The IFU states: use of 5% dextrose will result in 10% longer delivery time. D5W has different viscosity properties than the reference solution (0.9% NaCl at 32°C). Document and communicate to the care team if timing is clinically significant. |
| Q6 | Ans: C | The IFU explicitly contraindicates InfuLife for intra-articular infusion of anesthetics, reinforced in the Warnings section. All other options listed are approved indications or routes of administration. |
| Q7 | Ans: B | The IFU states: there is no alarm or alert when flow interruption occurs; therefore, InfuLife is not recommended for life-supporting medications. The outer cover scale provides only a general indication of remaining volume and is not accurate. |
| Q8 | Ans: B | Step 3 of Directions for Filling states: close the clamp and fill the device with no more than the recommended maximum fill volume. The clamp must be closed before filling to prevent uncontrolled flow from the distal end during preparation. |
| Q9 | Ans: B | The IFU states: nominal flow rates are based on sodium chloride (0.9%, 32°C) as reference, calculated per ISO 28620:2020, at 0 inches of head height. Any deviation — different fluid, temperature, or height — will alter actual flow from the nominal value. |
| Q10 | Ans: C | The IFU states under Residual Risks: InfuLife is contraindicated for neonates, infants, and children weighing less than 20 kg due to potential exposure to EO residuals. Ethylene oxide used in sterilization; residual amounts may pose risk to smaller/younger patients. |
| Q11 | Ans: A | The IFU states: filling the pump less than nominal volume generally results in lower flow rate. Conversely, filling above nominal generally results in faster flow rate. Elastomeric bladder pressure is influenced by how fully it is filled. Underfilling also delivers less total drug. |
| Q12 | Ans: B | Starting the Infusion step 3 states: verify that the clamp on the tubing is closed before attaching to the patient. This prevents an uncontrolled bolus on connection. The clamp is then opened (step 5) to begin the infusion. |

PMP-1104, Rev A  ©2025 First Medical Source (FMS), LLC  |  InfuLife IFU PMP-1100 Rev E  |  Competency Assessment